Protocol for Polishing and Mercury Plating of gold electrodes

Sand the gold electrode in glass or PEEK with 400 grit sandpaper

(or more coarse sandpaper if necessary, then use 400 grit).

1st polish – 15 micron (brown) with white cloth (coarse)

2nd polish – 6 micron (yellow) with white cloth (coarse)

3rd polish – 1 micron (blue) with brown cloth (soft)

4th polish – ¼ micron (gray) with brown cloth (soft)

Length of time of polishing depends on the electrode. Good sanding is essential. You should never need more than 1-2 minutes on the 3rd and 4th polish. If you use the wrong cloths for polishes, you will not get a good electrode.

Plate the polished working gold electrode for 240 sec at -0.1 V vs the reference electrode in degassed acidic mercuric nitrate solution [3.4 g Hg(NO3)2 in 100 mL H2O containing about 300 microliters of concentrated HNO3] using a normal three electrode cell attached to a DLK analyzer. Stir or bubble argon while plating. The pH must be less than 1.5 so the Hg(II) stays dissolved – or it will hydrolyze. Use some concentrated nitric acid to dissolve the mercuric nitrate if it precipitates over time. After plating you will see a shiny silvery surface on the gold due to the Hg plating.

Polarize the gold-amalgam electrode in 1 M sodium hydroxide for 90 seconds. Attach the gold amalgam electrode to the – terminal on 9 V battery and attach a Pt wire to the + terminal of a 9 volt battery. No DLK necessary here. Make sure that the battery reads 9 V with a voltmeter! Polarization is essential for field measurements– always do it!!

N.B. – Double check the sign on the 9 volt battery for these polarizations. You can see bubbles forming at the gold electrode if done correctly. If you put the gold on the + terminal of the battery, you will lose your Hg plate and have to plate and/or polish again!!

Lastly, run about 10-20 CV scans in oxygenated seawater from -0.1 to -1.8 V using the DLK analyzer. This procedure is a check to see if the electrode is stable (you will get good oxygen and peroxide waves) and ready for calibration.

Preparation of Ag/AgCl solid state reference electrode

Place the silver wire or electrode and a Pt electrode in 1 M sodium chloride solution or clean seawater. Attach the silver electrode to the + terminal on a 9 V battery (to oxidize the Ag wire to Ag+) and attach the Pt wire to the - terminal of a 9 volt battery for 20-30 seconds.  You will see a discoloration of the silver, which is due to the formation of the AgCl precipitate on the Ag surface. No DLK necessary here.