## Summary
WHOTS | Years | Calibration Location | 300 kHz - Spin Errors | [9 empty] | 600 kHz - Spin Errors | [9 empty] | 1200 kHz - Spin Errors
|  |  | S/N | Pre-deployment | [5 empty] | Post-Deployment |  |  | S/N | Pre-deployment | [5 empty] | Post-Deployment |  |  | S/N | Pre-deployment | [5 empty] | Post-Deployment
|  |  |  | Before Cal |  |  | After Cal | [6 empty] | Before Cal |  |  | After Cal | [6 empty] | Before Cal |  |  | After Cal
|  |  |  | mean | low | high | mean | low | high | mean | low | high |  | mean | low | high | mean | low | high | mean | low | high |  | mean | low | high | mean | low | high | mean | low | high
1 | 2004-2005 | Snug/Boat | 4891 | n/a |  |  | n/a |  |  | n/a
2 | 2005-2006 | Boat/Snug | 4891 | n/a |  |  | n/a |  |  | -3.06 | -4.65 | -0.7
3 | 2006-2007 | Snug | 7637 | -1.92 | -5.01 | 0.77 | n/a |  |  | -7.5 | -10.19 | -4.4
4 | 2007-2008 | Snug | 4891 | -8.96 | -10.9 | -6.55 | -0.38 | -0.9 | 0 | -3.22 | -5 | -1 | 1825 | n/a |  |  | n/a |  |  | n/a
5 | 2008-2009 | Snug | 7637 | -4.64 | -6.4 | -3 | -4.36 | -4.85 | -3.7 | damaged | [12 empty] | 2530 | -3.32 | -9.2 | -0.2 | -0.2 | -1 | 0.5 | -10.24 | 11.9 | -9
6 | 2009-2010 | Snug/UH | 4891 | -11.3 | -15.3 | -7.85 | -14.08 | -14.55 | -13.4 | -1.78 | -3 | 0 | 1825 | n/a |  |  | -10.23 | -11.85 | -9.1 | 0.67 | -1 | 2
7 | 2010-2011 | Snug/UH | 7637 | -12.37 | -13.55 | -10.5 | -8.66 | -9.85 | -7.45 | 0.99 | 0.2 | 1.7 | 13917 | -7.72 | -8.6 | -6.45 | -10.09 | -10.9 | -9.15 | damaged
8 | 2011-2012 | UH | 4891 | -2.39 | -4.2 | -0.5 | -0.39 | -0.5 | -0.2 | -0.66 | -2.8 | 1.1 | 1825 | -1.31 | -2.2 | 0.5 | 1.09 | 0.5 | 1.4 | -1.06 | -2.3 | 0
9 | 2012-2013 | UH | 7637 | -2.03 | -8.7 | 6.9 | 0.61 | -0.3 | 3 | 1.83 | 0.5 | 3.1 | 13917 | -2.57 | -4.7 | -0.1 | -0.02 | -0.5 | 0.2 | 2.16 | -0.8 | 4.5
10 | 2013-2014 | UH | 4891 | -0.77 | -3.9 | 1.7 | -0.6 | -0.9 | -0.2 | -0.49 | -3.05 | 1.1 | 1825 | -0.72 | -3.2 | 2 | -0.91 | -1.3 | -0.6 | -3.16 | -4.05 | -2.4
11 | 2014-2015 | UH | 7637 | 4.71 | 2.6 | 8 | 4.33 | 4 | 4.75 | 0.82 | -2.75 | 3.95 | 13917 | 3.5 | 1.2 | 5.95 | 3.62 | 3.3 | 4 | 1.73 | -1.3 | 4
12 | 2015-2016 | UH | 4891 | 0.62 | -2.8 | 3.6 | 1.74 | 1.2 | 2 | 1.15 | 0.5 | 1.63 | 1825 | -0.28 | -3.8 | 2.85 | 0.2 | 0.1 | 0.35 | 0.05 | -1.03 | 0.95
13 | 2016-2017 | UH | 7637 | 0.93 | -4.15 | 5.25 | -0.74 | -1.35 | -0.45 | in water |  |  | 13917 | 1.26 | -1.55 | 4.55 | -0.05 | -0.45 | 0.2 | in water
Notes
| 1 | spin performed at sea aboard ship; no post-cruise cal recorded, turned around for redeployment
| 2 | turned around for redeployment, no pre-cruise cal recorded
| 3 | separate pre-spin in cage
| 4 | 1825 prepared by WHOI
| 5 | 7637 battery blew up
| 6 | 1825 prepared by WHOI? (and post-cal file notes "heading bias has not been applied")
| 7 | Boat docked 60 m. away during precal for 7637
##
updated: 7/26/16

## RDI Errors
WHOTS-1 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 4891 | 3.02 | 0.76 | 3.78 | 0.22 | 3.10 | 12.89 ±0.52 | 0.19 ±0.31
| After Calibration | 4891 | 0.16 | 0.20 | 0.36 | 0.21 | 0.32 | 1.33 ±0.61 | 8.58 ±0.40
Post Deployment |  | 4891 | not recorded
WHOTS-2 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 4891 | not recorded
| After Calibration | 4891
Post Deployment |  | 4891 |  |  |  |  | 1.92
WHOTS-3 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 7637 | 3.06 | 0.23 | 3.29 | 0.18 | 3.07 | 0.63 ± 0.94 | 1.15 ± 0.33
| After Calibration | 7637 | 0.18 | 0.13 | 0.31 | 0.12 | 0.26 | 2.02 ± 0.33 | 1.06 ± 0.31
Post Deployment |  | 7637 | 2.85 | 0.26 | 3.11 | 0.22 | 2.88 | -0.55 ± 0.59 | -0.51 ± 0.32
WHOTS-4 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 1825 | prepared by WHOI
|  | 4891 | 4.32 | 0.18 | 4.50 | 0.17 | 4.29 | 1.07 ± 0.27 | 12.10 ± 0.29
| After Calibration | 1825 | prepared by WHOI
|  | 4891 | 0.31 | 0.44 | 0.75 | 0.17 | 0.61 | 0.16 ± 0.27 | -10.07 ± 0.42
Post Deployment |  | 1825 | prepared by WHOI
|  | 4891 | 1.34 | 0.85 | 2.19 | 0.15 | 1.72 | 0.25 ± 0.34 | -0.50 ± 0.25
WHOTS-5 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 2530 | 4.58 | 1.33 | 5.91 | 0.38 | 5.16 | -0.74 ± 0.73 | 1.77 ± 0.74
|  | 7637 | 1.16 | 0.07 | 1.68 | 0.13 | 1.61 | -0.16 ± 0.70 | 1.40 ± 0.74
| After Calibration | 2530 | 0.29 | 0.63 | 0.92 | 0.14 | 0.84 | 0.37 ± 0.72 | 2.84 ± 0.78
|  | 7637 | 0.24 | 0.33 | 0.56 | 0.23 | 0.49 | -1.88 ± 0.81 | -0.48 ± 0.81
Post Deployment |  | 2530 | 1.44 | 0.77 | 2.22 | 0.16 | 1.91 | -0.56 ± 0.28 | -0.56 ± 0.21
|  | 7637 | *Damaged
WHOTS-6 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 1825 | 4.25 | 0.64 | 4.9 | 0.11 | 4.34 | 0.55 ± 0.28 | -0.30 ± 0.31
|  | 4891 | 3.84 | 0.57 | 4.4 | 0.18 | 3.95 | 0.52 ± 0.40 | -0.97 ± 0.39
| After Calibration | 1825 | 0.18 | 0.93 | 1.11 | 0.76 | 1.03 | 1.78 ± 1.76 | 11.51 ± 0.66
|  | 4891 | 0.24 | 0.68 | 0.92 | 0.24 | 0.83 | -13.90 ± 0.43 | -2.36 ± 0.69
Post Deployment |  | 1825 | 0.87 | 0.26 | 1.13 | 0.2 | 0.99 | 0.98 ± 0.56 | 4.47 ± 0.59
|  | 4891 | 2.19 | 0.43 | 2.62 | 0.11 | 2.28 | 0.80 ± 0.487 | 2.09 ± 0.45
WHOTS-7 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 13917 | 0.35 | 0.18 | 0.53 | 0.13 | 0.42 | -0.75 ± 0.45 | 1.10 ± 0.46
|  | 7637 | 1.20 | 0.76 | 1.96 | 0.10 | 1.70 | -1.02 ± 0.39 | -0.50 ± 0.37
| After Calibration | 13917 | 0.26 | 0.49 | 0.76 | 0.25 | 0.64 | 15.90 ± 0.66 | 3.09 ± 0.81
|  | 7637 | 1.01 | 0.79 | 1.80 | 0.29 | 1.50 | -3.02 ± 0.63 | 17.02 ± 0.57
Post Deployment |  | 13917 | *Damaged
|  | 7637 | 0.14 | 0.46 | 0.60 | 0.16 | 0.54 | -0.15 ± 0.81 | 0.71 ± 0.85
WHOTS-8 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 1825 | 1.27 | 0.44 | 1.71 | 0.11 | 1.48 | -1.81 ± 0.38 | 0.12 ± 0.37
|  | 4891 | 1.96 | 0.16 | 2.11 | 0.09 | 1.97 | 0.34 ± 0.52 | 0.35 ± 0.34
| After Calibration | 1825 | 0.01 | 0.22 | 0.23 | 0.14 | 0.22 | -15.45 ± 0.38 | -0.35 ± 0.47
|  | 4891 | 0.35 | 0.48 | 0.83 | 0.33 | 0.7 | -15.98 ± 0.56 | 0.08 ± 0.60
Post Deployment |  | 1825 | 1.13 | 0.19 | 1.33 | 0.14 | 1.15 | -1.21 ± 0.95 | 0.57 ± 0.79
|  | 4891 | 1.71 | 0.48 | 2.20 | 0.18 | 1.75 | 0.78 ± 1.04 | 0.46 ± 0.88
WHOTS-9 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 13917 | 2.43 | 0.29 | 2.73 | 0.13 | 2.49 | 1.45 ± 0.87 | -0.29 ± 0.79
|  | 7637 | 8.19 | 0.68 | 8.88 | 0.17 | 8.23 | 1.00 ± 0.85 | -0.74 ± 0.8
| After Calibration | 13917 | 0.18 | 0.1 | 0.28 | 0.12 | 0.23 | 12.29 ± 0.79 | -0.29 ± 0.83
|  | 7637 | 0.45 | 0.12 | 0.57 | 0.20 | 0.48 | -14.71 ± 0.89 | -0.23 ± 0.82
Post Deployment |  | 13917 | 2.85 | 0.12 | 2.97 | 0.18 | 2.87 | -0.30 ± 0.74 | 0.80 ± 1.03
|  | 7637 | 0.95 | 0.51 | 1.46 | 0.19 | 1.22 | 0.90 ± 0.76 | -0.31 ± 0.73
WHOTS-10 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 1825 | 2.81 | 0.05 | 2.87 | 0.13 | 2.81 | -0.13 ± 0.82 | -1.44 ± 1.05
|  | 4891 | 2.82 | 0.70 | 3.52 | 0.12 | 3.08 | 1.10 ± 0.89 | -1.22 ± 0.85
| After Calibration | 1825 | 0.03 | 0.15 | 0.19 | 0.15 | 0.18 | 0.35 ± 1.05 | 14.20 ± 0.96
|  | 4891 | 0.22 | 0.20 | 0.41 | 0.12 | 0.35 | -17.09 ± 0.90 | -1.23 ± 0.90
Post Deployment |  | 1825 | 0.73 | 0.11 | 0.84 | 0.27 | 0.76 | -2.48 ± 0.89 | -0.43 ± 0.75
|  | 4891 | 1.97 | 0.32 | 2.29 | 0.21 | 2.06 | -0.01 ± 0.20 | 0.82 ± 0.73
WHOTS-11 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 13917 | 2.80 | 0.45 | 3.25 | 0.17 | 2.84 | 1.26 ± 0.89 | -0.77 ± 0.87
|  | 7637 | 2.29 | 0.16 | 2.45 | 0.22 | 2.29 | 0.28 ± 1.16 | -0.43 ± 1.04
| After Calibration | 13917 | 0.35 | 0.31 | 0.66 | 0.14 | 0.54 | -16.86 ± 0.73 | -0.75 ± 0.80
|  | 7637 | 0.42 | 0.27 | 0.68 | 0.13 | 0.59 | -2.01 ± 0.71 | -0.29 ± 0.71
Post Deployment |  | 13917 | 2.73 | 0.21 | 2.94 | 0.18 | 2.76 | -1.63 ± 1.39 | 0.09 ± 1.30
|  | 7637 | 3.38 | 0.34 | 3.72 | 0.18 | 3.45 | 0.07 ± 1.58 | -0.36 ± 1.18
WHOTS-12 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 1825 | 3.26 | 0.06 | 3.32 | 0.17 | 3.25 | 2.04 ± 0.28 | 0.21 ± 0.27
|  | 4891 | 2.98 | 0.27 | 3.25 | 0.18 | 3.01 | 1.17 ± 0.25 | 0.50 ± 0.25
| After Calibration | 1825 | 0.18 | 0.13 | 0.32 | 0.14 | 0.28 | -16.58 ± 0.34 | 0.24 ± 0.36
|  | 4891 | 0.27 | 0.2 | 0.47 | 0.21 | 0.38 | 13.59 ± 0.38 | 0.02 ± 0.34
Post Deployment |  | 1825 | 1 | 0.09 | 1.09 | 0.3 | 1 | 1.04 ± 0.19 | 1.36 ± 0.63
|  | 4891 | 0.54 | 0.29 | 0.84 | 0.21 | 0.65 | -0.26 ± 0.13 | 0.73 ± 0.63
WHOTS-13 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 13917 | 3.05 | 0.18 | 3.23 | 0.14 | 3.06 | -0.41 ± 0.62 | 0.76 ± 0.72
|  | 7637 | 4.87 | 0.41 | 5.28 | 0.17 | 4.94 | 0.09 ± 0.59 | -1.40 ± 0.60
| After Calibration | 13917 | 0.14 | 0.37 | 0.51 | 0.24 | 0.47 | -0.06 ± 0.64 | 17.75 ± 0.67
|  | 7637 | 0.66 | 0.02 | 0.68 | 0.26 | 0.66 | 0.51 ± 0.60 | 16.97 ± 0.68
Post Deployment |  | 13917
7637

## Comp
s/n | WHOTS | Pre Deployment |  |  |  |  | Pre Deployment |  |  |  |  | Post Deployment
|  | Before Cal |  |  |  |  | After Cal
|  | @ | Spin Data Errors |  |  | RDI Estimate | @ | Spin Data Errors |  |  | RDI Estimate | @ | Spin Data Errors |  |  | RDI Estimate
|  |  | Mean | Min-Mean | Max-Mean | Overall Error(±) |  | Mean | Min-Mean | Max-Mean | Overall Error(±) |  | Mean | Min-Mean | Max-Mean | Overall Error(±)
4891 | 1 | Snug | n/a |  |  | 3.1 | Snug | n/a |  |  | 0.32 | Sea | n/a |  |  | n/a
| 2 | Sea | n/a |  |  | n/a | Sea | n/a |  |  | n/a | Snug | -3.06 | -1.59 | 2.36 | 1.92
| 4 | Snug | -8.96 | -1.94 | 2.41 | 4.29 | Snug | -0.38 | -0.52 | 0.38 | 0.61 | Snug | -3.22 | -1.78 | 2.22 | 1.72
| 6 | Snug | -11.3 | -4 | 3.45 | 3.95 | Snug | -14.08 | -0.47 | 0.68 | 0.83 | UH | -1.78 | -1.22 | 1.78 | 2.28
| 8 | UH | -2.39 | -1.81 | 1.89 | 1.97 | UH | -0.39 | -0.11 | 0.19 | 0.7 | UH | -0.66 | -2.14 | 1.76 | 1.75
| 10 | UH | -0.77 | -3.13 | 2.47 | 3.08 | UH | -0.6 | -0.3 | 0.4 | 0.35 | UH | -0.49 | -2.56 | 1.59 | 2.06
| 12 | UH | 0.62 | -3.42 | 2.98 | 3.01 | UH | 1.74 | -0.54 | 0.26 | 0.38 | UH | 1.15 | -0.65 | 0.48 | 0.65
Mean |  | Snug | -10.13 | ± | 2.95 | 3.78 | Snug | -7.23 | ± | 0.5125 | 0.58667 | Snug | -3.14 | ± | 1.9875 | 1.82
Mean |  | UH | -0.84667 | ± | 2.6167 | 2.6867 | UH | 0.25 | ± | 0.3 | 0.47667 | UH | -0.445 | ± | 1.5225 | 1.685
Mean |  | ALL | -4.56 | ± | 2.75 | 3.2333 | ALL | -2.742 | ± | 0.385 | 0.53167 | ALL | -1.3433 | ± | 1.6775 | 1.73
s/n | WHOTS | Pre Deployment |  |  |  |  | Pre Deployment |  |  |  |  | Post Deployment
|  | Before Cal |  |  |  |  | After Cal
|  | @ | Spin Data Errors |  |  | RDI Estimate | @ | Spin Data Errors |  |  | RDI Estimate | @ | Spin Data Errors |  |  | RDI Estimate
|  |  | Mean | Min-Mean | Max-Mean | Overall Error(±) |  | Mean | Min-Mean | Max-Mean | Overall Error(±) |  | Mean | Min-Mean | Max-Mean | Overall Error(±)
7637 | 3 | Snug | -1.92 | -3.09 | 2.69 | 3.07 | Snug | n/a |  |  | 0.26 | Snug | -7.5 | -2.69 | 3.1 | 2.88
| 5 | Snug | -4.64 | -1.76 | 1.64 | 1.61 | Snug | -4.36 | -0.49 | 0.66 | 0.49 | Snug | damaged
| 7 | Snug | -12.37 | -1.18 | 1.87 | 1.7 | Snug | -8.66 | -1.19 | 1.21 | 1.5 | UH | 0.99 | -0.79 | 0.71 | 0.54
| 9 | UH | -2.03 | -6.67 | 8.93 | 8.23 | UH | 0.61 | -0.91 | 2.39 | 0.48 | UH | 1.83 | -1.33 | 1.27 | 1.22
| 11 | UH | 4.71 | -2.11 | 3.29 | 2.29 | UH | 4.33 | -0.33 | 0.42 | 0.59 | UH | 0.82 | -3.57 | 3.13 | 3.45
| 13 | UH | 0.93 | -5.08 | 4.32 | 4.94 | UH | -0.74 | -0.61 | 0.29 | 0.66 | UH | in water
Mean |  | Snug | -6.31 | ± | 2.0383 | 2.1267 | Snug | -6.51 | ± | 0.8875 | 0.75 | Snug | -7.5 | ± | 2.895 | 2.88
Mean |  | UH | 1.2033 | ± | 5.0667 | 5.1533 | UH | 1.4 | ± | 0.825 | 0.57667 | UH | 1.2133 | ± | 1.8 | 1.7367
Mean |  | ALL | -2.5533 | ± | 3.5525 | 3.64 | ALL | -1.764 | ± | 0.85 | 0.66333 | ALL | -0.965 | ± | 2.0737 | 2.0225
s/n | WHOTS | Pre Deployment |  |  |  |  | Pre Deployment |  |  |  |  | Post Deployment
|  | Before Cal |  |  |  |  | After Cal
|  | @ | Spin Data Errors |  |  | RDI Estimate | @ | Spin Data Errors |  |  | RDI Estimate | @ | Spin Data Errors |  |  | RDI Estimate
|  |  | Mean | Min-Mean | Max-Mean | Overall Error(±) |  | Mean | Min-Mean | Max-Mean | Overall Error(±) |  | Mean | Min-Mean | Max-Mean | Overall Error(±)
1825 | 4 | Snug | whoi |  |  |  | Snug | whoi |  |  |  | Snug | whoi
| 6 | Snug | n/a |  |  | 4.34 | Snug | -10.23 | -1.62 | 1.13 | 1.03 | UH | 0.67 | -1.67 | 1.33 | 0.99
| 8 | UH | -1.31 | -0.89 | 1.81 | 1.48 | UH | 1.09 | -0.59 | 0.31 | 0.22 | UH | -1.06 | -1.24 | 1.06 | 1.15
| 10 | UH | -0.72 | -2.48 | 2.72 | 2.81 | UH | -0.91 | -0.39 | 0.31 | 0.18 | UH | -3.16 | -0.89 | 0.76 | 0.76
| 12 | UH | -0.28 | -3.52 | 3.13 | 3.25 | UH | 0.2 | -0.1 | 0.15 | 0.28 | UH | 0.05 | -1.08 | 0.9 | 1
Mean |  | Snug | n/a |  |  | 4.34 | Snug | -10.23 | ± | 1.375 | 1.03 | Snug | n/a
Mean |  | UH | -0.77 | ± | 2.425 | 2.5133 | UH | 0.12667 | ± | 0.30833 | 0.22667 | UH | -0.875 | ± | 1.1162 | 0.975
Mean |  | ALL | -0.77 | ± | 2.425 | 2.97 | ALL | -3.35 | ± | 0.575 | 0.4275 | ALL | -0.875 | ± | 1.1162 | 0.975
s/n | WHOTS | Pre Deployment |  |  |  |  | Pre Deployment |  |  |  |  | Post Deployment
|  | Before Cal |  |  |  |  | After Cal
|  | @ | Spin Data Errors |  |  | RDI Estimate | @ | Spin Data Errors |  |  | RDI Estimate | @ | Spin Data Errors |  |  | RDI Estimate
|  |  | Mean | Min-Mean | Max-Mean | Overall Error(±) |  | Mean | Min-Mean | Max-Mean | Overall Error(±) |  | Mean | Min-Mean | Max-Mean | Overall Error(±)
13917 | 7 | Snug | -7.72 | -0.88 | 1.27 | 0.42 | Snug | -10.09 | -0.81 | 0.94 | 0.64 | UH | damaged
| 9 | UH | -2.57 | -2.13 | 2.47 | 2.49 | UH | -0.02 | -0.48 | 0.22 | 0.23 | UH | 2.16 | -2.96 | 2.34 | 2.87
| 11 | UH | 3.5 | -2.3 | 2.45 | 2.84 | UH | 3.62 | -0.32 | 0.38 | 0.54 | UH | 1.73 | -3.03 | 2.27 | 2.76
| 13 | UH | 1.26 | -2.81 | 3.29 | 3.06 | UH | -0.05 | -0.4 | 0.25 | 0.47 | UH | in water
Mean |  | Snug | -7.72 | ± | 1.075 | 0.42 | Snug | -10.09 | ± | 0.875 | 0.64 | Snug | n/a
Mean |  | UH | 0.73 | ± | 2.575 | 2.7967 | UH | 1.1833 | ± | 0.34167 | 0.41333 | UH | 1.945 | ± | 2.65 | 2.815
Mean |  | ALL | -1.3825 | ± | 1.9556 | 2.2025 | ALL | -1.635 | ± | 0.42222 | 0.47 | ALL | 1.945 | ± | 2.65 | 2.815
s/n | WHOTS | Pre Deployment |  |  |  |  | Pre Deployment |  |  |  |  | Post Deployment
|  | Before Cal |  |  |  |  | After Cal
|  | @ | Spin Data Errors |  |  | RDI Estimate | @ | Spin Data Errors |  |  | RDI Estimate | @ | Spin Data Errors |  |  | RDI Estimate
|  |  | Mean | Min-Mean | Max-Mean | Overall Error(±) |  | Mean | Min-Mean | Max-Mean | Overall Error(±) |  | Mean | Min-Mean | Max-Mean | Overall Error(±)
2530 | 5 | Snug | -3.32 | -5.88 | 3.12 | 5.16 | Snug | -0.2 | -0.8 | 0.7 | 0.84 | Snug | -10.24 | -1.66 | 1.24 | 1.91
Mean |  | Snug | -3.32 | ± | 4.5 | 5.16 | Snug | -0.2 | ± | 0.75 | 0.84 | Snug | -10.24 | ± | 1.45 | 1.91

## sn4891
WHOTS | Years | Calibration Location | 300 kHz - Spin Errors
|  |  | S/N | Pre-deployment | [5 empty] | Post-Deployment
|  |  |  | Before Cal |  |  | After Cal
|  |  |  | mean | low | high | mean | low | high | mean | low | high
1 | 2004-2005 | Snug/Boat | 4891 | n/a |  |  | n/a |  |  | n/a
2 | 2005-2006 | Boat/Snug | 4891 | n/a |  |  | n/a |  |  | -3.06 | -4.65 | -0.7
4 | 2007-2008 | Snug | 4891 | -8.96 | -10.9 | -6.55 | -0.38 | -0.9 | 0 | -3.22 | -5 | -1
6 | 2009-2010 | Snug/UH | 4891 | -11.3 | -15.3 | -7.85 | -14.08 | -14.55 | -13.4 | -1.78 | -3 | 0
8 | 2011-2012 | UH | 4891 | -2.39 | -4.2 | -0.5 | -0.39 | -0.5 | -0.2 | -0.66 | -2.8 | 1.1
10 | 2013-2014 | UH | 4891 | -0.77 | -3.9 | 1.7 | -0.6 | -0.9 | -0.2 | -0.49 | -3.05 | 1.1
12 | 2015-2016 | UH | 4891 | 0.62 | -2.8 | 3.6 | 1.74 | 1.2 | 2 | 1.15 | 0.5 | 1.63
WHOTS-1 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 4891 | 3.02 | 0.76 | 3.78 | 0.22 | 3.10 | 12.89 ±0.52 | 0.19 ±0.31
| After Calibration | 4891 | 0.16 | 0.20 | 0.36 | 0.21 | 0.32 | 1.33 ±0.61 | 8.58 ±0.40
Post Deployment |  | 4891 | not recorded
WHOTS-2 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 4891 | not recorded
| After Calibration | 4891
Post Deployment |  | 4891 |  |  |  |  | 1.92
WHOTS-4 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 4891 | 4.32 | 0.18 | 4.50 | 0.17 | 4.29 | 1.07 ± 0.27 | 12.10 ± 0.29
| After Calibration | 4891 | 0.31 | 0.44 | 0.75 | 0.17 | 0.61 | 0.16 ± 0.27 | -10.07 ± 0.42
Post Deployment |  | 4891 | 1.34 | 0.85 | 2.19 | 0.15 | 1.72 | 0.25 ± 0.34 | -0.50 ± 0.25
WHOTS-6 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 4891 | 3.84 | 0.57 | 4.4 | 0.18 | 3.95 | 0.52 ± 0.40 | -0.97 ± 0.39
| After Calibration | 4891 | 0.24 | 0.68 | 0.92 | 0.24 | 0.83 | -13.90 ± 0.43 | -2.36 ± 0.69
Post Deployment |  | 4891 | 2.19 | 0.43 | 2.62 | 0.11 | 2.28 | 0.80 ± 0.487 | 2.09 ± 0.45
WHOTS-8 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 4891 | 1.96 | 0.16 | 2.11 | 0.09 | 1.97 | 0.34 ± 0.52 | 0.35 ± 0.34
| After Calibration | 4891 | 0.35 | 0.48 | 0.83 | 0.33 | 0.7 | -15.98 ± 0.56 | 0.08 ± 0.60
Post Deployment |  | 4891 | 1.71 | 0.48 | 2.20 | 0.18 | 1.75 | 0.78 ± 1.04 | 0.46 ± 0.88
WHOTS-10 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 4891 | 2.82 | 0.70 | 3.52 | 0.12 | 3.08 | 1.10 ± 0.89 | -1.22 ± 0.85
| After Calibration | 4891 | 0.22 | 0.20 | 0.41 | 0.12 | 0.35 | -17.09 ± 0.90 | -1.23 ± 0.90
Post Deployment |  | 4891 | 1.97 | 0.32 | 2.29 | 0.21 | 2.06 | -0.01 ± 0.20 | 0.82 ± 0.73
WHOTS-12 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 4891 | 2.98 | 0.27 | 3.25 | 0.18 | 3.01 | 1.17 ± 0.25 | 0.50 ± 0.25
| After Calibration | 4891 | 0.27 | 0.2 | 0.47 | 0.21 | 0.38 | 13.59 ± 0.38 | 0.02 ± 0.34
Post Deployment |  | 4891 | 0.54 | 0.29 | 0.84 | 0.21 | 0.65 | -0.26 ± 0.13 | 0.73 ± 0.63

## sn7637
WHOTS | Years | Calibration Location | 300 kHz - Spin Errors
|  |  | S/N | Pre-deployment | [5 empty] | Post-Deployment
|  |  |  | Before Cal |  |  | After Cal
|  |  |  | mean | low | high | mean | low | high | mean | low | high
3 | 2006-2007 | Snug | 7637 | -1.92 | -5.01 | 0.77 | n/a |  |  | -7.5 | -10.19 | -4.4
5 | 2008-2009 | Snug | 7637 | -4.64 | -6.4 | -3 | -4.36 | -4.85 | -3.7 | damaged
7 | 2010-2011 | Snug/UH | 7637 | -12.37 | -13.55 | -10.5 | -8.66 | -9.85 | -7.45 | 0.99 | 0.2 | 1.7
9 | 2012-2013 | UH | 7637 | -2.03 | -8.7 | 6.9 | 0.61 | -0.3 | 3 | 1.83 | 0.5 | 3.1
11 | 2014-2015 | UH | 7637 | 4.71 | 2.6 | 8 | 4.33 | 4 | 4.75 | 0.82 | -2.75 | 3.95
13 | 2016-2017 | UH | 7637 | 0.93 | -4.15 | 5.25 | -0.74 | -1.35 | -0.45 | in water
WHOTS-3 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 7637 | 3.06 | 0.23 | 3.29 | 0.18 | 3.07 | 0.63 ± 0.94 | 1.15 ± 0.33
| After Calibration | 7637 | 0.18 | 0.13 | 0.31 | 0.12 | 0.26 | 2.02 ± 0.33 | 1.06 ± 0.31
Post Deployment |  | 7637 | 2.85 | 0.26 | 3.11 | 0.22 | 2.88 | -0.55 ± 0.59 | -0.51 ± 0.32
WHOTS-5 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 7637 | 1.16 | 0.07 | 1.68 | 0.13 | 1.61 | -0.16 ± 0.70 | 1.40 ± 0.74
| After Calibration | 7637 | 0.24 | 0.33 | 0.56 | 0.23 | 0.49 | -1.88 ± 0.81 | -0.48 ± 0.81
Post Deployment |  | 7637 | *Damaged
WHOTS-7 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 7637 | 1.20 | 0.76 | 1.96 | 0.10 | 1.70 | -1.02 ± 0.39 | -0.50 ± 0.37
| After Calibration | 7637 | 1.01 | 0.79 | 1.80 | 0.29 | 1.50 | -3.02 ± 0.63 | 17.02 ± 0.57
Post Deployment |  | 7637 | 0.14 | 0.46 | 0.60 | 0.16 | 0.54 | -0.15 ± 0.81 | 0.71 ± 0.85
WHOTS-9 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 7637 | 8.19 | 0.68 | 8.88 | 0.17 | 8.23 | 1.00 ± 0.85 | -0.74 ± 0.8
| After Calibration | 7637 | 0.45 | 0.12 | 0.57 | 0.20 | 0.48 | -14.71 ± 0.89 | -0.23 ± 0.82
Post Deployment |  | 7637 | 0.95 | 0.51 | 1.46 | 0.19 | 1.22 | 0.90 ± 0.76 | -0.31 ± 0.73
WHOTS-11 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 7637 | 2.29 | 0.16 | 2.45 | 0.22 | 2.29 | 0.28 ± 1.16 | -0.43 ± 1.04
| After Calibration | 7637 | 0.42 | 0.27 | 0.68 | 0.13 | 0.59 | -2.01 ± 0.71 | -0.29 ± 0.71
Post Deployment |  | 7637 | 3.38 | 0.34 | 3.72 | 0.18 | 3.45 | 0.07 ± 1.58 | -0.36 ± 1.18
WHOTS-13 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 7637 | 4.87 | 0.41 | 5.28 | 0.17 | 4.94 | 0.09 ± 0.59 | -1.40 ± 0.60
| After Calibration | 7637 | 0.66 | 0.02 | 0.68 | 0.26 | 0.66 | 0.51 ± 0.60 | 16.97 ± 0.68
Post Deployment |  | 7637

## sn1825
WHOTS | Years | Calibration Location | 600 kHz - Spin Errors
|  |  | S/N | Pre-deployment | [5 empty] | Post-Deployment
|  |  |  | Before Cal |  |  | After Cal
|  |  |  | mean | low | high | mean | low | high | mean | low | high
4 | 2007-2008 | Snug | 1825 | n/a |  |  | n/a |  |  | n/a
6 | 2009-2010 | Snug/UH | 1825 | n/a |  |  | -10.23 | -11.85 | -9.1 | 0.67 | -1 | 2
8 | 2011-2012 | UH | 1825 | -1.31 | -2.2 | 0.5 | 1.09 | 0.5 | 1.4 | -1.06 | -2.3 | 0
10 | 2013-2014 | UH | 1825 | -0.72 | -3.2 | 2 | -0.91 | -1.3 | -0.6 | -3.16 | -4.05 | -2.4
12 | 2015-2016 | UH | 1825 | -0.28 | -3.8 | 2.85 | 0.2 | 0.1 | 0.35 | 0.05 | -1.03 | 0.95
WHOTS-4 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 1825 | prepared by WHOI
| After Calibration | 1825 | prepared by WHOI
Post Deployment |  | 1825 | prepared by WHOI
WHOTS-6 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 1825 | 4.25 | 0.64 | 4.9 | 0.11 | 4.34 | 0.55 ± 0.28 | -0.30 ± 0.31
| After Calibration | 1825 | 0.18 | 0.93 | 1.11 | 0.76 | 1.03 | 1.78 ± 1.76 | 11.51 ± 0.66
Post Deployment |  | 1825 | 0.87 | 0.26 | 1.13 | 0.2 | 0.99 | 0.98 ± 0.56 | 4.47 ± 0.59
WHOTS-8 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 1825 | 1.27 | 0.44 | 1.71 | 0.11 | 1.48 | -1.81 ± 0.38 | 0.12 ± 0.37
| After Calibration | 1825 | 0.01 | 0.22 | 0.23 | 0.14 | 0.22 | -15.45 ± 0.38 | -0.35 ± 0.47
Post Deployment |  | 1825 | 1.13 | 0.19 | 1.33 | 0.14 | 1.15 | -1.21 ± 0.95 | 0.57 ± 0.79
WHOTS-10 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 1825 | 2.81 | 0.05 | 2.87 | 0.13 | 2.81 | -0.13 ± 0.82 | -1.44 ± 1.05
| After Calibration | 1825 | 0.03 | 0.15 | 0.19 | 0.15 | 0.18 | 0.35 ± 1.05 | 14.20 ± 0.96
Post Deployment |  | 1825 | 0.73 | 0.11 | 0.84 | 0.27 | 0.76 | -2.48 ± 0.89 | -0.43 ± 0.75
WHOTS-12 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 1825 | 3.26 | 0.06 | 3.32 | 0.17 | 3.25 | 2.04 ± 0.28 | 0.21 ± 0.27
| After Calibration | 1825 | 0.18 | 0.13 | 0.32 | 0.14 | 0.28 | -16.58 ± 0.34 | 0.24 ± 0.36
Post Deployment |  | 1825 | 1 | 0.09 | 1.09 | 0.3 | 1 | 1.04 ± 0.19 | 1.36 ± 0.63

## sn13917
WHOTS | Years | Calibration Location | 600 kHz - Spin Errors
|  |  | S/N | Pre-deployment | [5 empty] | Post-Deployment
|  |  |  | Before Cal |  |  | After Cal
|  |  |  | mean | low | high | mean | low | high | mean | low | high
7 | 2010-2011 | Snug/UH | 13917 | -7.72 | -8.6 | -6.45 | -10.09 | -10.9 | -9.15 | damaged
9 | 2012-2013 | UH | 13917 | -2.57 | -4.7 | -0.1 | -0.02 | -0.5 | 0.2 | 2.16 | -0.8 | 4.5
11 | 2014-2015 | UH | 13917 | 3.5 | 1.2 | 5.95 | 3.62 | 3.3 | 4 | 1.73 | -1.3 | 4
13 | 2016-2017 | UH | 13917 | 1.26 | -1.55 | 4.55 | -0.05 | -0.45 | 0.2 | in water
WHOTS-7 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 13917 | 0.35 | 0.18 | 0.53 | 0.13 | 0.42 | -0.75 ± 0.45 | 1.10 ± 0.46
| After Calibration | 13917 | 0.26 | 0.49 | 0.76 | 0.25 | 0.64 | 15.90 ± 0.66 | 3.09 ± 0.81
Post Deployment |  | 13917 | *Damaged
WHOTS-9 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 13917 | 2.43 | 0.29 | 2.73 | 0.13 | 2.49 | 1.45 ± 0.87 | -0.29 ± 0.79
| After Calibration | 13917 | 0.18 | 0.1 | 0.28 | 0.12 | 0.23 | 12.29 ± 0.79 | -0.29 ± 0.83
Post Deployment |  | 13917 | 2.85 | 0.12 | 2.97 | 0.18 | 2.87 | -0.30 ± 0.74 | 0.80 ± 1.03
WHOTS-11 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 13917 | 2.80 | 0.45 | 3.25 | 0.17 | 2.84 | 1.26 ± 0.89 | -0.77 ± 0.87
| After Calibration | 13917 | 0.35 | 0.31 | 0.66 | 0.14 | 0.54 | -16.86 ± 0.73 | -0.75 ± 0.80
Post Deployment |  | 13917 | 2.73 | 0.21 | 2.94 | 0.18 | 2.76 | -1.63 ± 1.39 | 0.09 ± 1.30
WHOTS-13 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 13917 | 3.05 | 0.18 | 3.23 | 0.14 | 3.06 | -0.41 ± 0.62 | 0.76 ± 0.72
| After Calibration | 13917 | 0.14 | 0.37 | 0.51 | 0.24 | 0.47 | -0.06 ± 0.64 | 17.75 ± 0.67
Post Deployment |  | 13917

## sn2530
WHOTS | Years | Calibration Location | 1200 kHz - Spin Errors
|  |  | S/N | Pre-deployment | [5 empty] | Post-Deployment
|  |  |  | Before Cal |  |  | After Cal
|  |  |  | mean | low | high | mean | low | high | mean | low | high
5 | 2008-2009 | Snug | 2530 | -3.32 | -9.2 | -0.2 | -0.2 | -1 | 0.5 | -10.24 | -11.9 | -9
WHOTS-5 |  | ADCP (s/n) | Single Cycle Error (°) | Double Cycle Error (°) | Largest Double + Single Cycle Error (°) | RMS of 3rd Order and Higher + Random Error (°) | Overall Error (°) | Pitch, Mean and St. Dev. (°) | Roll, Mean and St. Dev. (°)
Pre Deployment | Before Calibration | 2530 | 4.58 | 1.33 | 5.91 | 0.38 | 5.16 | -0.74 ± 0.73 | 1.77 ± 0.74
| After Calibration | 2530 | 0.29 | 0.63 | 0.92 | 0.14 | 0.84 | 0.37 ± 0.72 | 2.84 ± 0.78
Post Deployment |  | 2530 | 1.44 | 0.77 | 2.22 | 0.16 | 1.91 | -0.56 ± 0.28 | -0.56 ± 0.21