SBE    Sea-Bird Electronics, Inc.

1808 136th  Place N.E., Bellevue, Washington 98005 USA

Telephone: (425) 643-9866   Telex: 292915 SBEI UR    Fax: (425) 643-9954

## PRESSURE TEST CERTIFICATE

Date:	24 Aug 01

Job #	26851 R

Model	9plus 6800m

S/N	09P8235-0310

Pressure test results:

Low pressure (psi)	50	psi

Time (minutes)	15	min

High pressure (psi)	1000	psi

Time (minutes)	30	min

Pass               Yes

# Fail	_________

Comments:

Replaced the housing and the 6 pin connectors at JT2 and JT7.

Tested by:       CPE

-- -- - - - - - - - - - - - - - - - - - - - - - -

Pressure Calibration Check			22 August 2001

pressure sensor model: Digiquartz 410K-105

sensor serial number:    51421

installed in:                   CTD 09P8235-0310

This pressure calibration is a check of the 'test' sensor against a stable reference pressure sensor.  The reference pressure sensor is itself checked several times per year against a NIST-traceable pressure standard maintained at Paroscientific, Inc. . The circumstances of this pressure check introduce no more than 1.5 psia total error in 10,000 psi (0.015 %) in addition to the error resident in the Paroscientific site standard.  The check offers a very high level certification of the health and proper operation of the 'test' sensor.

Input                     Sensor             Sensor         Pressure          Pressure              Error

Pressure*              Output         Temperature  Factory Coef    Corrected

[psia]	[hz]	[deg C]	[psia]	[psia]	[psia]

14.700	33007.58	23.0	18.177	15.314	0.614

2014.709	33835.67	23.1	2017.380	2014.358	-0.351

4014.447	34641.21	23.2	4017.485	4014.304	-0.143

6013 .761	35425.20	23.2	6017.027	6013.687	-0.074

8013.178	36189.36	23.3	8016.790	8013.292	0.114

10012.926	36934.88	23.3	10016.504	10012.847	-0.079

8013.205	36189.40	23.4	8016.956	8013.458	0.253

6013.749	35425.23	23.4	6017.217	6013.878	0.129

4014.247	34641.09	23.4	4017.338	4014.157	-0.090

2014.555	33835.50	23.5	2017.157	2014.135	-0.420

14.557	33007.17	23.5	17.468	14.605	0.048

Input pressure is generated with a Ruska model 5201 dead-weight tester, serial number 23330/380, and is determined by measurement with reference pressure sensor model Digiquartz 410K-000, serial number 73292.

Sensor Temperature: pressure sensor internal temperature.

Pressure Corrected: pressure computed with original factory coefficients and then corrected with a slope and offset to give the best linear agreement with the ‘reference’ Input pressure.

Error: Corrected pressure - Input pressure

A linear fit of this calibration data, between sensor pressure computed with factory coefficients and the Input pressure, yields correction coefficients:

Corrected pressure = psi_slope * Factory pressure + psi_offset [psia]

psi_slope = 0.99992 and psi_offset = -2.86 [psia]

These are converted to Slope and Offset in decibars for use in the SEASOFT programs by:

Slope = psi_slope = 0.99992

Offset = C * (psi_offset - 14.7 * (1 - psi_slope)) = -1.9739 [dbars]

C = 0.689476 [dbar/psi]

Slope and Offset coefficients are entered into the pressure sensor calibration coefficient section of the <>.CON file using the program SEACON.

Digiquartz Coefficients:	AD590 Pressure Temperature Coefficients:

C1	=	-3.928393e+04

C2	=	1.087860e+00	AD590M =	0.01151

C3	=	1.199240e-02	AD590B =	-8.74988

D1	=	3.882500e-02

D2	=	0.000000e+00

T1	=	3.029977e+01	Calibration	Correction:

T2	=	4.275780e-05

T3	=	4.335010e-06	Slope = 0.99992

T4	=	2.205920e-09	Offset = -1.9739